## Slide 1: Centralized vacuum or suction system
By: Dr. Azad J. Ali

## Slide 2: Suction Systems
It plays a crucial role in many parts as :
Operating theatre.
Anesthesia room (induction room).
Recovery room.
Intensive Care Unit (ICU).
Patient’s room( ward).

## Slide 4: Components
1-A pump or a power source (-500 mmHg)
2-A suction controller with a filter.
3-A receiver or a collection vessel.
4-A suction tubing and suction nozzle
(e.g. Yankauer nozzle) or catheter.

## Slide 5: Central suction system
Pump or Vacuum generator
Vacuum controller, filter & receiver(collection vessel)

## Slide 6: Suction tubing transparent , flexible , firm & of sufficient diameter

## Slide 7: Yankauer Nozzle

## Slide 8: Suction catheter
Different sizes

## Slide 9: Mechanism of Action
Negative pressure generator (Electrical or Pneumatic or using venturi principle).
Manual control of the amount of –ve pressure.
Variable 0riffice with a float assembly.
Receiver of sufficient capacity.
Suction tubing:
-transparent.
-flexible& firm.
-sufficient diameter.

## Slide 10: Negative pressure generator
Electrical
Pneumatic (Mechanical)

## Slide 11: Illustration of vacuum system

## Slide 12: Illustration of Suction apparatus

## Slide 13
Venturi Negative pressure generator

## Slide 14: Testing the efficacy ( check-up) of central-piped Vacuum systems before use
A negative pressure of at least(-400 mmHg) to be maintained at the outlet.
Withstand a free flow of at least (40L/min) at the outlet.
Should take no longer than 10 sec. to generate a vacuum of (-500 mmHg) with (25L/min) air displacement.

## Slide 15: Thank you